# Closed-loop experiments with CLoSES-SEM

## Up Front Summary

- Get code from:

  - CLoSES-SEM: Trial by Trial Neural Decoder Closed-loop system (to stimulate following decoded cognitive state)

(https://github.com/Center-For-Neurotechnology/CLoSES-SEM.git)

- Steps:

  - Intracranial electrical brain signals from NSPs or acquisition boards are acquired,

  - Bipolar or Referential montage,

  - Band pass filtered,

  - Features (power or coherence) continuously computed once a trial starts.

  - After epochs specified in model (default 2s), features are averaged and used as input to decoder model.

  - Decoder model runs and produces mean / boundaries estimates

  - If state is above (below) thresholds, stimulation is triggered on the following trial (default).

- Key characteristics:

  - Features are computed per trial, averaged over time epochs, and a neural encoder/decoder model estimates a hidden state.

  - If decoded state is above (below) threshold, stimulation occurs on following trial.

- For safety, only N out of M trials can result on stimulation.

- It runs on a dedicated Simulink Real-Time computer with a GUI on another computer for configuration and real-time visualization.

- It runs in parallel to clinical and research systems.

- Replay of previously acquired datasets allows feature selection and model training.

# Hardware & Connections in the EMU

## Non-Stim day hardware & connections – with BlackRock NSP acquisition

- USB labeled CLoSES (connected to CLoSES computer in patient’s room).

- If tasks:

  - Presentation computer with task

  - Connect image onset trigger (CHANNEL 2 from Presentation Octopus) from presentation computer to NSP AINP3 (corresponds to channel 131 on CLoSES GUI).

- Connect channel 1 of octopus from CLoSES to AINP2 (to have “Stimulation” on NSX file)

- Connect channel 2-3 of octopus from CLoSES to AINP5-6 (to have “Detections” events on NSX file)

- Connect channel 5-8 of octopus from CLoSES to AINP7-10 (to have “stim channel” on NSX file)

- All connections should be split to NSP2.

## Stim day hardware & connections– with BlackRock NSP acquisition

- All of Non-Stim hardware & connections (above)

- Connect an octopus cable to parallel port cable from CLoSES computer (labelled CLoSES).

  - Connect channel 1 of octopus from the CLoSES to TRIGGER of Cerestim (this Triggers the Stimulation).

  - This BNC (Cerestim TRIGGER) should also be connected to NSP AINP2.

  - Connect the other channels to the other AINP inputs

- Stimulator (Cerestim)

  - Check connections (Headbox, power cord, and blue cables).

  - USB from Cerestim connected to Presentation computer (different for ECR!)

  - SYNC output connected to AINP1

  - CHECK that TRIGGER is connected to Octopus channel 1 from CLoSES and to AINP2

- Optional: Connect output from Behavioral GUI to NSP AINP4 (corresponds to channel 132 on CLoSES GUI).

# Running CLoSES-SEM in the EMU

Configuring CLoSES and GUI:

- To Start CLoSES:

  - Connect USB key (black with label: CLoSES) to CLoSES target computer.

  - Turn on CLoSES target computer. (Black screen with “Simulink” on the bottom should appear on small monitor).

- On HOST computer (BlackRock Computer - #1):

  - Open MATLAB 2016b

  - Go to …\MATLAB\CLoSES\CLoSES-SEM (see below)

- In runClosedLoopStateEstimate.m edit participant name

- Run runClosedLoopStateEstimate.m

  - Alternatively:

- Add CLoSES-SEM to your path (right click on folder CLoSES-SEM and select Add to Path with SubFolders).

  - ClosedLoopWithNeuralModel should be on path. If it is not:

  - Type: GUINeuralDecoderClosedLoopConsole (‘PatientName’,’MGxxx’)

- RECOMMENDED change directory where data is saved: GUINeuralDecoderClosedLoopConsole('PatientName','test','dirResults','D:\DeciderData')

# Running CLoSES-SEM in Replay

CLoSES-RT can run on your Windows computer to replay data previously acquired.

Running CLoSES-SEM:

  - Open MATLAB 2016b

  - Go to …\MATLAB\CLoSES\CLoSES-SEM (see below)

- In runClosedLoopStateEstimate.m edit participant name

- Run runClosedLoopStateEstimate.m

  - Alternatively:

- Add CLoSES-SEM to your path (right click on folder CLoSES-SEM and select Add to Path with SubFolders).

  - ClosedLoopWithNeuralModel should be on path. If it is not:

  - Type: GUINeuralDecoderClosedLoopConsole (‘PatientName’,’MGxxx’)

- RECOMMENDED change directory where data is saved: GUINeuralDecoderClosedLoopConsole('PatientName','test','dirResults','D:\DeciderData')

# Operating the CLoSES-SEM GUI

On Initialization screen:

  - Select EXPERIMENT TYPE on drop-down menu (e.g. MSIT)

  - Load Neural Model: select your model file from your directory (Usually on .\ExampleData)

  - OPTIONAL:

  - If you want to load pre-determined channels or other parameters:

- Click on Load Config and look for configuration file (Default ones are on folder .\Configuration)

- ConfigurationFile_MSIT_FixThreshold.m for FIXED threshold

- Click on Load Model and look for MODEL file (C:\Users\User\Documents\MATLAB\ ClosedLoopWithNeuralModel\ CompileFiles\ ClosedLoopStimXpcTarget_WithDecoderModel.dlm)

- NOTE: If you are on directory C:\Users\User\Documents\MATLAB\ ClosedLoopWithNeuralModel no need to “load model”

- Click on Closed-Loop (to load main GUI)

On CLoSES-SEM GUI:

  - On Channels Panel: Select contacts (bipolar channels correspond to left-right pairs of contacts in order (e.g. 1-2; 2-3 etc.)

- NOTE: NSP2 channels start at channel 201 (e.g. NSP channel 14 corresponds to channel 214 on CLoSES GUI)

EXAMPLE:

- BLOCK 1: Right selector: 232,233,234Left selector: 233,234,235

  - Right dACC- Channels RMF1-RMF2, channel numbers-:  232-233(NSP 2)

  - Right dACC- Channels RMF2-RMF3, channel numbers-: 233-234(NSP 2)

  - Right dACC- Channels RMF3-RMF4, channel numbers-: 234-235(NSP 2)

- BLOCK 2: Right selector: 27,28,29 Left selector: 28,29,30

  - Left dACC- Channels LMF1-LMF2, channel numbers-: 27-28 (NSP 1)

  - Left dACC- Channels LMF2-LMF3, channel numbers-: 28-29 (NSP 1)

  - Left dACC- Channels LMF3-LMF4, channel numbers-: 29-30 (NSP 1)

- By Default, ALL features are selected as “Detection Channels/Pairs” to save all features (only those on dValid are processed by decoder model).

- OPTIONAL: Select as Visualization channels only those in dValid (only those used by model)

- Other Options should be double checked (see also figure below):

EXAMPLE: for MSIT with stimulation if Upper bound is larger than Threshold:

- Filter: THETAALPHAGAMMA

- Feature: LOGBandPower

- Detector: NeuralModel

- Stim on: Next Trial

- State Output: UpperBound

- Click on Compile to Compile model (it will take 1-2min)

- Click on Start to start Closed-Loop Stimulation System

- Click on Stop to stop and save file

        - There are 2 files: on for continuous data (EEG and stim) and one for trial by trial data (features, thresholds, state).

        - Continuous Data File is called: DeciderData_EXPERIMENTTYPE_PATIENTNAME_DATE_TIME.mat

(e.g. DeciderData_MSIT_test_200301_0846.mat)

        - Trial by Trial Data File is called: DeciderTrialByTrialData_EXPERIMENTTYPE_PATIENTNAME_DATE_TIME.mat      	(e.g. DeciderTrialByTrialData_MSIT_test_200301_0846.mat)

# When Running CLoSES-SEM GUI

- Once started, you can change the threshold value and use SendParam to change the levels.

      - Since MSIT only uses 1 threshold remove “ThBelow” by unchecking checkbox

- Stop – Start after each block to get separate files.

      - REMEMBER: Only saves data to files when STOP is pressed.

NOTE: this GUI won’t do anything unless you have the task running, so start the task!

Note when stim is happening, that it is occurring with trials and trial-based features.

## Recover Data

- Data will be saved on Folder D:\DeciderData (or folder specified in ‘dirResults’,’DIRNAME’).

- Get the code from the components that generated the tasks and set up so that the code could be checked as to what it did per patient later

- Make sure to copy every file with today’s date on them (there will be 1 file per block)

  - Make sure to Stop/Start for each block

# Installing CLoSES-SEM

Get Git bash: http://msysgit.github.io/

Go to the Github TRANSFORM website: https://github.com/Center-For-Neurotechnology/CLoSES-SEM.git

Copy repository information to Clipboardhttps://github.com/Center-For-Neurotechnology/CLoSES-SEM.git

To get the code, go to the directory you want to install things, clone the project.

Step by step:

- Open GitHub bash (http://msysgit.github.io/ or terminal on Linux/MAC)

- Go to the directory where CLoSES will be installed (cd /D/MATLAB/CLoSES/CLoSES-SEM).

- Type: git clone https://github.com/Center-For-Neurotechnology/CLoSES-SEM.git  .

(the last dot means here).

## Updating CLoSES-SEM code:

To update the code, open git bash, go to the directory you installed things and type: git pull (see example below).

The first time you will be asked to relate your copy to origin. Follow instructions on screen.

If you are updating on a computer for simulations, that’s it!

# Compiling model for the first time

If you are updating the RIG, it is a good idea to compile model before running GUI (see screenshot below).

- In MATLAB type: 						initializationScriptNeuralModel('REAL-TIME','ExampleData/identityModel_15_1.mat')

- Double-click on .\ClosedLoopStimulation\StateEstimateModel\ClosedLoopStimXpcTarget_WithDecoderModel.slx

- Go to directory: CompileFiles

- On model click the “Build model”

- Get back to CLoSES-SEM root folder

# Dependencies

MATLAB toolboxes:

- MATLAB

- Simulink

- Simulink Real-Time

- Signal Processing Toolbox

- Communications Toolbox

- DSP System Toolbox